Cold Chain Management Policy and Procedure Manual

INTRODUCTION TO COLD CHAIN MANAGEMENT	2

What is the cold chain?	2

Why is cold chain management important?	2

OUR COLD CHAIN MANAGEMENT PROFILE	2

Our people	2

Our vaccine refrigerator	3

Our monitoring equipment	3

ORDERING, RECEIVING AND DISPOSAL OF VACCINES	4

Ordering vaccines	4

Receiving vaccines	5

Disposal of vaccines	5

PACKING OUR VACCINE REFRIGERATOR	5

Principles of packing our vaccine refrigerator	5

Packing our purpose-built vaccine refrigerator	5

CHECKING AND RECORDING THE VACCINE REFRIGERATOR TEMPERATURE	6

Principles of monitoring our vaccine refrigerator	6

Monitoring vaccines using a data logger	7

MANAGEMENT OF COLD CHAIN PROBLEMS	7

Planning and prevention	7

Managing a power failure with a purpose-built vaccine refrigerator	8

Action in the event of a cold chain breach	9

PORTABLE COOLERS	10

Monitoring a cooler	10

Equipment for a portable cooler	11

Packing a portable cooler	11

Conditioning ice packs/gel packs	11

MAINTAINING OUR EQUIPMENT	12

Maintaining our vaccine refrigerator	12

Location of our vaccine refrigerator	12

Maintaining our monitoring equipment	13

Knowing the accuracy of the external thermometer (“slush” test)	13

INTRODUCTION TO COLD CHAIN MANAGEMENT

What is the cold chain?

The National Vaccine Storage Guidelines: Strive for 5 define the cold chain as the system of transporting and storing vaccines within the safe temperature range of +2°C and +8°C. The cold chain begins from the time the vaccine is manufactured, moves through to the state or territory vaccine distribution centres and ends when the vaccine is administered. The success of the system involves three key elements: people, processes and equipment.

Why is cold chain management important?

Vaccines are delicate biological substances that can become less effective or destroyed if they are frozen, allowed to get too hot and/or exposed to direct sunlight or fluorescent light. The outcome following exposure to any or all of these conditions is dependent on the vaccine itself. For example, for vaccines that are cold or freeze-sensitive such as pneumococcal and influenza vaccines, the loss of potency following freezing is immediate and these vaccines must not be administered. When vaccines are exposed to repeated episodes of heat, the loss of vaccine potency is cumulative and cannot be reversed.

OUR COLD CHAIN MANAGEMENT PROFILE

Our people

In our practice, the following staff members are responsible for specified activities:

| Activity | Staff Details |
| --- | --- |
| Primary person responsible for cold chain management | Practice manager/nurse – Coco Gao |
| Secondary person responsible for cold chain management | Practice nurse – Scarlet Liang |
| Recording the temperature in the morning | Practice nurse/receptionist during the weekend |
| Recording the temperature in the evening | Practice nurse/ receptionist during the weekend |
| Ordering vaccines | Practice nurse – Coco Gao |
| Receiving vaccines | Practice nurse – Coco Gao, Scarlet Liang and Bonnie Shang |
| Checking vaccine expiry dates and rotating stock | Practice nurse – Coco Gao, Scarlet Liang and Bonnie Shang |

All new members of the practice team, particularly those who are involved in some or all elements of cold chain management, are provided with a comprehensive orientation program to ensure they are well-versed in their role and responsibilities, and are well-informed of the person(s) with primary responsibility for cold chain management.

All key staff members receive regular updates on cold chain management by attending education sessions, reading newsletter and journal articles or from engagement and support offered by our Primary Health Network, local Public Health Unit, professional associations (e.g. APNA, AAPM, RACGP, ACRRM), accreditation bodies (AGPAL), and other like organisations.

Our vaccine refrigerator

In our practice, we use the following dedicated refrigerator to store vaccines:

| Refrigerator model | Avem Quirks Medisafe Plus FKG311 |
| --- | --- |
| Name of supplier | Avem Quirks |
| Contact details of supplier | 1800 658 120 |
| Warranty expiry date | Expired |
| Name of maintenance service provider | Avem Quirks |
| Contact details of maintenance service provider | 1800 658 120/1800 804 627 |

Our vaccine refrigerator offers the following features to help us achieve effective cold chain management:

Abnormal temperature alarm system for low and high temperatures

Refrigerator door closes automatically

Built-in power filters to protect from power surges

Refrigerator is linked into our practice email and phone, which will send immediate notifications if the refrigerator experiences a power failure or is outside the recommended temperature range of +2°C and +8°C

Audible alarm activates during power failure

Two backup batteries which can provide up to 20 hours run time during power failure

Our monitoring equipment

In our practice, we use the following equipment to monitor the temperature of our vaccine refrigerator:

Internal minimum/maximum thermometer

Internal Data logger: The AQ Box

Back-up external minimum/maximum thermometer

We record the temperature of our vaccine refrigerator using the procedure listed under the Checking and recording the vaccine refrigerator temperature section within this document.

We maintain our equipment using the procedure listed under the Maintaining our monitoring equipment section within this document.

ORDERING, RECEIVING AND DISPOSAL OF VACCINES

Ordering vaccines

At our practice, we keep vaccine stock to a minimum by regularly monitoring the vaccines stored in the vaccine refrigerator.

We order our private vaccines on a monthly basis through the following suppliers:

| Vaccine | Supplier | Contact Details |
| --- | --- | --- |
| Private vaccines | Dorevitch Pathology | 03 9244 0230 |
|  | GSK | 1800 077 421 |
|  | Seqirus | 03 9389 2000 |
|  | EBOS | 1800 269 534 |
|  | Team Medical | 1300 22 44 50 |
| Government vaccines | Onelink | 1800 014 207 |
| Covid 19 vaccines | CVAS | 1800 318 208 |

We order our publicly funded vaccines through the Victorian Department of Health and human services on a monthly basis by email via orders@onlink.com.au.

When placing an order, the Victorian Government Department of Health representative will ask:

What is your immunisation service provider number?

Have your vaccines been stored between minimum of +2°C and maximum of +8°C since your last order?

What quantities of each vaccine are remaining? (i.e. stock on hand)

What is the expiry date of each vaccine?

A distributor will deliver the ordered vaccines to our practice in refrigerated delivery trucks packed in a cooler with freezer bricks, freeze indicator and dual time- temperature indicator.

Receiving vaccines

In our practice, all members of the practice team are educated on the urgency of undertaking the processes associated with receiving vaccines. Vaccines are received by the appointed personnel who are responsible for:

Signing the delivery docket for the courier.

Checking that the consignment is correct against the delivery invoice, and that the vaccines are packed appropriately.

Unpacking the vaccines into the vaccine refrigerator as soon as possible.

Checking the temperature indicators contained within the ordered stock to ensure the vaccines have not been compromised.

heat indicator shows colour change in the centre circle of the indicator

freeze indicators have colour bulbs that release a dye at or below 0°C.

If there are any concerns about the vaccines delivered, the vaccines are isolated in the vaccine refrigerator and contact is made with Victorian Government Department of Health as soon as possible after receiving the delivery.  For private vaccines, isolate the affected stock and notify the supplier immediately.

Disposal of vaccines

Discarding vaccines occurs as a result of expiration or a consequence to a cold chain breach.

Publicly funded vaccines are only discarded on advice from Victorian Government Department of Health.

According to the Poisons Standard June 2017 (www.tga.gov.au/publication/poisons-standard-susmp), vaccines are classed as Schedule 4 and are therefore disposed in accordance with our practice procedure on clinical waste management.

PACKING OUR VACCINE REFRIGERATOR

Principles of packing our vaccine refrigerator

Correct packing of our vaccine refrigerator is essential if vaccines are to remain safe and effective. The vaccine refrigerator needs to have capacity to accommodate the maximum vaccine storage needs, including during the influenza vaccination period.

Packing our purpose-built vaccine refrigerator

In our practice, we follow these steps when packing our purpose-built vaccine refrigerator:

Vaccines are stored in their original packaging from the manufacturer - this helps to protect them from temperature fluctuations and ultraviolet (UV) light.

Vaccines are stored in shallow plastic baskets/trays clearly labelled with the name of vaccine, and adequate space between each basket/tray to allow for air circulation.

Vaccines are rotated so that vaccines with the shortest expiry dates are used first.

There is a gap of at least 4cm between the vaccines and the back of the refrigerator.

When storing influenza vaccines, vaccines for adults and vaccines children under five (5) years of age are separated, clearly labelled and stored in separate areas of the refrigerator.

Privately purchased vaccines are clearly marked and separated from the National Immunisation Program vaccines.

If there is a small amount of stock in the refrigerator, cooled bottles of water or refrigerated ice packs/gel packs are placed in the unused space to help stabilise the temperature.

The refrigerator is used exclusively for the storage of vaccines. Medications requiring refrigeration may also be stored providing there is sufficient space. It is not used to store food or drink.

A sticker is placed on the refrigerator door reminding GPs and staff to only open the door when required.

A warning sticker is placed above the refrigerator power source stating “DO NOT TURN POWER OFF”.

CHECKING AND RECORDING THE VACCINE REFRIGERATOR TEMPERATURE

Principles of monitoring our vaccine refrigerator

Checking and recording the minimum and maximum temperature of the vaccine refrigerator is an essential element of ensuring that vaccines remain safe and effective. Checking and recording temperatures before using vaccines enables the identification of problems before a vaccine is given.

Twice daily temperature checking gives a better indication of any problems in the refrigerator’s function and temperature fluctuations over the course of the day; however, the temperature needs to be viewed and considered every time the refrigerator is opened.

In addition to the twice daily monitoring and recording of the vaccine refrigerator temperature when our practice is open, our practice ensures safe and effective vaccines by:

Checking minimum/maximum temperatures

on receipt of vaccines

hourly during outreach clinics,

following a power failure

every time the refrigerator door is opened.

Recording in the temperature log comments and actions taken when the temperature is recorded outside the recommended range of +2°C and +8°C (e.g. restocking).

Resetting the thermometer only after temperatures have been recorded.

Maintaining a separate temperature log for each vaccine refrigerator in the practice.

Monitoring vaccines using a data logger

In our practice, we follow this procedure for checking and recording the temperature of our vaccine refrigerator:

- At the beginning of the first session (i.e. at the start of the business day)

Download the readings of the data logger.

Check both the minimum (MIN) and maximum (MAX) temperatures.

Record the date, time, the minimum and maximum temperature and your initials on the temperature chart or log.

At the end of the last session (i.e. at the close of the business day)

Download the readings of the data logger.

Check both the minimum (MIN) and maximum (MAX) temperatures.

Record the date, time, the minimum and maximum temperature and your initials on the temperature chart or log.

If the vaccine refrigerator temperatures have been recorded outside the recommended range of +2°C and +8°C, follow the process outlined in Action in the event of a cold chain breach.

MANAGEMENT OF COLD CHAIN PROBLEMS

Planning and prevention

Cold chain problems can take many shapes and forms such as inadequate equipment and processes, incorrect packing of the vaccine refrigerator, inadequate training for staff involved in cold chain management, vaccines inadvertently left out of the refrigerator, refrigerator accidentally turned off or unplugged, power failure, load shedding of electricity, and significant weather events such as cyclones and floods. All of these events can lead to a cold chain breach.

In our practice, we have put the following measures in place to prevent, wherever possible, a cold chain breach:

Simple, routine processes established.

Written protocols developed on safe and effective cold chain management.

Staff involved in cold chain management receive adequate training and understand the cold chain and its importance.

Appropriate vaccine refrigerator and monitoring equipment is utilised.

Processes are in place for ordering and receiving vaccines.

Processes are in place for safe disposal of vaccines.

The vaccine refrigerator is packed in accordance with best practice guidelines.

The temperature of the vaccines is monitored and recorded in accordance with best practice guidelines.

Processes are in place for managing a power failure.

Equipment and processes are in place for packing a portable cooler.

Processes are in place for maintenance of the vaccine refrigerator and of the monitoring equipment.

A cold chain management self-audit is conducted on an annual basis.

Cold chain risk management planning is undertaken.

A review of processes in the event that a cold chain problem occurs.

Managing a power failure with a purpose-built vaccine refrigerator

Purpose-built vaccine refrigerators (particularly those with glass doors) may lose their chill quicker than a domestic refrigerator, often in as a little as 20 to 30 minutes. According to our refrigerator and back up battery manufacturer, our practice’s purpose-built vaccine refrigerator linking with UPS back up batteries is able to maintain a temperature of +2°C and +8°C in the event of a power failure for up to 20 hours

In our practice, we follow this procedure in the event of a power failure during business hours:

- Investigate the reason for the power failure.

If it is a power cut, telephone the utility company to ascertain approximately how long the power will be interrupted.

If a safety switch (Residual Current Device) has tripped, reset it. If it trips again, contact an electrician.

- Frequently monitor the remaining time of the UPS back up battery.

- Frequently monitor the temperature of the refrigerator.

- As purpose-built vaccine refrigerators can warm quickly once back up batteries run flat, add cooled water bottles or refrigerated ice packs/gel packs to the vaccine refrigerator to help keep it cool.

- Use an alternative means of vaccine storage (such as a cooler, frozen ice packs/gel packs and insulating material) where required.

- Where the vaccines are transferred to a portable cooler, continue to monitor the temperature of the vaccines by placing the thermometer probe inside a vaccine box inside the cooler. Monitoring occurs every 15 minutes for the first two (2) hours as freezing is most likely to occur during this period. Following the initial two hour period, the cooler is monitored every hour. Refer to the processes outlined in Packing a portable cooler contained in this document.

IMPORTANT: Depending on the circumstances of a power failure, ice packs/gel packs may not be given adequate conditioning time prior to packing a portable cooler. In these instances, additional insulating material is used to protect the vaccines, and the portable cooler monitored closely.

Action in the event of a cold chain breach

The National Vaccine Storage Guidelines: Strive for 5 defines a cold chain breach as exposure of vaccines to temperatures outside the recommended range of +2°C to +8°C (excluding fluctuations up to +12°C, lasting no longer than 15 minutes, when stock taking or restocking).

Cold chain breaches left unidentified and untreated can have serious implications – especially when it involves informing people that they, or their child, may have received an ineffective vaccine and will require re-vaccination.

In our practice, we follow this procedure in the event of a cold chain breach:

- Immediately isolate the vaccines and label ‘Do not use‘.

- Keep vaccines refrigerated between +2°C and +8°C.

- Contact the Victorian Government Department of Health as soon as possible during business hours and have important details on hand including:

The practice’s immunisation service provider number

Date and time of the breach

The type of refrigerator in which the vaccines are stored (i.e. purpose-built or domestic)

Download the data logger to indicate the temperature range during the breach period

Minimum and maximum temperature reading

Are cold chain monitors (CCMs) stored with the vaccines? If ‘yes’, be ready to report the reading when the breach was noticed

Length of time the refrigerator temperature was outside +2°C and +8°C

Position of the temperature probe and vaccines in the refrigerator

Has the vaccine refrigerator had any maintenance issues recently?

Period of time that these problems have been occurring

Type and number of vaccines in current stock

Expiry date of the vaccines

Have any vaccines been pushed up against the cooling plate or a cold air outlet?

Are all vaccines in their original packaging?

What was the cause of the cold chain breach and has it been rectified?

Has anybody been vaccinated with potentially affected vaccines?

Have the vaccines previously been exposed to temperatures outside of the +2°C and +8°C range?

- Do not discard any vaccine until advised to do so by the Victorian Government Department of Health. For privately purchased vaccines, contact the manufacturer for advice.

- Record notes on the temperature log regarding what happened and how the problem was corrected.

PORTABLE COOLERS

A cooler, also known by names such as Esky™ or Willow™, is a solid-walled insulated container with a tight fitting lid with the temperature inside maintained by ice packs or gel packs.

Although generally associated with outreach immunisation clinics, our practice has portable coolers in a range of sizes available for the following circumstances:

Transport of vaccines.

Defrosting a domestic vaccine refrigerator.

During a power failure.

Vaccine refrigerator breakdown.

Vaccine refrigerator not maintaining the correct temperatures of between +2°C and +8°C.

Other circumstances requiring the practice to remove vaccines from the refrigerator (e.g. significant weather event such as a cyclone).

It is important to note that freezing happens very easily in all coolers, usually in the first two (2) hours after packing, and therefore it is important that the cooler selected for use in any of the above situations is large enough to store the vaccine as well as sufficient insulating material.

IMPORTANT: In the event of a significant weather event such as a cyclone, or depending on the circumstances of a power failure, ice packs/gel packs may not be given adequate conditioning time prior to packing a portable cooler. In these instances, we use additional insulating material to protect the vaccines, and we monitor the portable cooler closely.

Monitoring a cooler

In our practice, we check the temperature of the cooler:

After packing.

Every 15 minutes for the first two (2) hours, and then every hour following.

Prior to administering vaccine.

Before returning the vaccines to the vaccine refrigerator.

Equipment for a portable cooler

In our practice, we have the following equipment available to pack a portable cooler:

Coolers (such as an Esky™ or Willow™) in sizes that meet the practice needs, and are all a minimum of 10 litres.

Ice packs and/or gel packs.

Insulating material

polystyrene sheets (width 12-20mm)

polystyrene chips

plastic bubble-wrap

shredded paper.

Minimum/maximum thermometer with probe.

NOTE: As our practice uses a purpose-built vaccine refrigerator, the ice packs/gel packs for the cooler are stored in the freezer compartment of the domestic refrigerator in the staff lunchroom.

Packing a portable cooler

- In our practice, we follow this procedure to pack a portable cooler:

- Chill the inside of the cooler prior to use by placing ice packs/gel packs in it for a few hours.

- Place insulating material at the bottom of the container.

- Use a minimum/maximum thermometer to monitor the temperature inside the cooler. The probe is placed inside an empty vaccine box with the product information leaflet.

- Surround the vaccines with insulating material.

- When using a small cooler, place the conditioned ice packs/gel packs on top, close and seal the lid of the cooler.

- When using a large portable cooler, place conditioned ice packs/gel packs around the sides of the cooler as well as on top.

- Ensure vaccines are not in direct contact with the ice packs/gel packs to minimise risk of freezing.

Conditioning ice packs/gel packs

Conditioning means leaving the ice packs/gel packs at room temperature to allow the ice or gel at the core to rise to about 0°C. This is also known as ‘sweating’.

Ice packs/gel packs must be conditioned correctly before use as the risk of freezing vaccines increases if the ice packs/gel packs are not conditioned correctly. It is noted, however, that in the event of a significant weather event (such as a cyclone) or depending on the circumstances of a power failure, ice packs/gel packs may not be given adequate conditioning time prior to packing a portable cooler. In these instances, we use additional insulating material to protect the vaccines, and monitor the portable cooler closely.

In our practice, in addition to following the manufacturer’s instructions on correct conditioning of ice packs/gel packs, we follow this procedure:

- Remove ice packs/gel packs from the freezer.

- Lay out in a single row on their sides.

- Leave a 5cm space around each ice pack/gel pack to allow maximum air exposure to reduce conditioning time.

Conditioning time will depend on the ambient temperature, and the type of ice pack/gel pack selected including the size/weight of the ice pack/gel pack.

MAINTAINING OUR EQUIPMENT

A well maintained vaccine refrigerator and monitoring equipment is essential if vaccines are to remain safe and effective.

For vaccine refrigerators, it involves having a maintenance program in place and positioning the refrigerator in a suitable location.

Ensuring monitoring equipment is effective involves changing the battery every year and doing an annual check of the monitoring equipment to identify potentially faulty equipment, flat batteries, or a damaged probe or cable which can affect the degree of accuracy of the temperature readings.

Maintaining our vaccine refrigerator

In our practice, we maintain our vaccine refrigerator in the following ways:

Conduct a vaccine management audit (including people, processes and equipment) at least every 12 months.

Report refrigerator problems immediately to the nominated personnel so that appropriate action, including repairs where necessary, can commence.

Check the rubber seal around the door

If brittle or torn, arrange for replacement.

As there are exposed coils on the back of the refrigerator, we keep the coils clean and dust free to improve operating efficiency.

Location of our vaccine refrigerator

In our practice, we have positioned our vaccine refrigerator according to these principles:

Place the vaccine refrigerator away from warm external walls and out of direct sunlight.

Ensure the refrigerator is in a secure area only accessible to staff.

Follow the manufacturer’s instructions about positioning the refrigerator to enable sufficient air circulation around the back and sides.

Ensure the power source is labelled clearly to prevent the refrigerator from being accidentally unplugged or turned off. Where the power source is exposed, we use a switch cover.

Maintaining our monitoring equipment

In our practice, we maintain our monitoring equipment by doing the following:

Replace the battery of the data logger at least every 12 months (or sooner if thermometer or data logger issues are experienced).

Conduct an accuracy check of the external thermometer (also called “slush” test) when:

Commissioning a new thermometer.

Changing the thermometer battery.

At least every 12 months (or sooner where thermometer or cold chain issues are experienced). Record the results on the temperature chart or log for future reference.

Record the actions and results of the above when undertaken, on the temperature log.

Knowing the accuracy of the external thermometer (“slush” test)

In our practice, we follow this procedure to know the accuracy of the external thermometer:

- Two-thirds fill a polystyrene or plastic cup with cold water.

- Place cup in any available refrigerator freezer in the practice until a fine layer of ice forms on the top and small sections of ice form within the fluid. This may take up to 2 ½ hours.

Using this method, the mixture is 0°C if ice is present.

- Place the temperature probe into the middle of the container. Be careful not to let the probe touch the container.

- Observe the temperature on the display screen after two (2) minutes has lapsed.

The temperature will drop quickly at first and then more slowly; however, the temperature should drop to 0°C within two (2) minutes.

Interpreting the results:

- The results should be within one (1) degree above or below 0°C, therefore the display screen may show three possible readings: +1°C, 0°C, -1°C.

- If the temperature reading is more than one degree above or below 0°C at two (2) minutes, replace the battery and test again.

- If still not within range, contact Victorian Government Department of Health for further advice.

- If the temperature reading is 0°C this demonstrates that the thermometer is accurate and can be used as normal.

- If the temperature reading is either +1°C or -1°C this demonstrates that the accuracy is within acceptable limits and can continue to be used; however, such results need to be:

Recorded on the temperature log.

Taken into account should the thermometer record temperatures outside the +2°C to +8°C range.

Included in the details given to Victorian Government Department of Health when reporting a cold chain breach.